EndoFit Stentgraft - Endovascular Repair of TAA : Short and Midterm Results of a Single Center

Luigi Inglese1, M.D., Nadia Mollichelli1,   M.D., Massimo Medda1, M.D, Cristina Sirolla4, M.D., Valerio Tolva2, M.D., Viviana Grassi2, M.D., Cecilia Fantoni1, M.D., Marco Pavesi3, M.D.

1 Cardiovascular Interventional laboratoratory, San Donato Milanese Hospital, San Donato Milanese (Milano), Italy

2    Vascular Surgery, San Donato Milanese Hospital, San Donato Milanese (Milano), Italy

3   Division of Anesthesiology, San Donato Milanese Hospital, San Donato Milanese (Milano), Italy

4    Biostatistical Centre, Dept. Of Gerontological Research, INRCA

Address for reprints and correspondence:

Prof. Luigi Inglese

Policlinico San Donato Milanese

Via Morandi 30, 20097 San Donato Milanese (Milano), Italy

Tel: +39. 02.52774388

Fax: +39. 02.52774585

E-mail: luigi.inglese@tin.it

Abstract

Background: Compared with open repair, endovascular treatment (EVT) for Thoracic Aortic Aneurysm (TAA) has been shown to reduce the mortality rates at endovascular centers globally. Mortality rates are obviously the primary concern of any TAA endovascular specialist, while other established benefits of TAA Endografts are: reduced OR time, less blood loss, diminished ICU patient stay and limited major perioperative/postoperative complications. The aim of this study is to evaluate and analyse the outcomes of the endovascular treatment of thoracic aortic pathology performed at a single center with the use of the LeMaitre Vascular EndoFit stent graft system, manufactured by LeMaitre Vascular, Burlington, Ma. USA.

Methods: From January 2002 to January 2007, 41 patients were treated with the EndoFit stent graft system for thoracic aortic disease. All patients’ data were obtained through a retrospective review of hospital records. Indications for treatment were progression of aneurysm size in atherosclerotic aneurysms (n=24, mean aneurismal diameter: 7.191.48), acute contained aortic rupture (n=5, two of these presenting aortoesophageal fistula as a complicating factor), recurrent pain and persistent hypertension or progression of aneurysm size despite maximal medical therapy in patients with aortic dissection (n=6), recurrent pain in patients with penetrating atherosclerotic ulcers (n=4), post-traumatic pseudoaneurysm ( n=1) and post-decoartation aneurysm (n=1).

Results: In all 41 patients, the EndoFit Stent Graft was successfully deployed. The in hospital and thirty day mortality rate was 7.32% (three patients). We observed three postoperative endoleaks (7.32%): one was a proximal type I endoleak (Ia), the second was a distal (Ib) endoleak, both of which were self-repaired during one and three months follow up period. The third patient was discovered with a type Ia endoleak after one year and it was necessary to perform an open surgical conversion. During follow up, two secondary type Ia endoleaks were observed and treated with an additional stent graft. Acute myocardial infarction, need for dialysis, respiratory failure, graft migration, infection or fracture were not observed in our patients. There was only one spinal ischemia with consequent lower extremity weakness. No paraplegia was observed in our patients. There were seven deaths (17%) in total during the follow up period, but only one was aneurysm related. Average follow up was 24.818 months. At two years, overall patients survival by Kaplan-Meier analysis was 70% and aneurysm related survival was 89%.

Conclusion: Endovascular treatment of vascular disease involving the descending thoracic aorta can be safely performed with the EndoFit Thoracic Stent Graft system. We observed few vascular access complications, no fracture or migration of the stent-graft, only one spinal ischemia and an aneurysm related survival rate of 89% at two years.

Key words: EndoFit, endovascular repair, thoracic aortic disease.

Introduction

The incidence and prevalence of thoracic aortic aneurysm are universally increasing. This can be attributed to the aging of the population and the increased use of diagnostic procedures in other thoracic pathologies 1, 2.  For specific pathologies as atherosclerotic aneurysm, dissection, intramural hematoma, penetrating ulcer, or post-traumatic pseudo-aneurysm, open surgical repair of thoracic aorta has been considered the gold standard treatment strategy 3, 4. However, endovascular repair of thoracic aortic disease has proven to reduce in comparison to open surgery perioperative mortality, operative time, ICU stay and major complications, especially spinal cord ischemia with resultant paraplegia5, 6.  The risks of open surgery of thoracic aorta are prohibitive in an elderly population (>75 years), mainly in those having multiple co-morbidities (cardiac, pulmonary or renal). Endovascular treatment of thoracic aortic disease may offer to selected patients with suitable aortic anatomy both short and midterm advantages to open surgery, as it avoids a thoracotomy and thoracic aortic clamping.

From October 1999 to January 2007, our Institution 96 patients were treated with thoracic endografts, acquiring experience and analysing the characteristics of the different and most commonly used endografts: Talent (AVE/Medtronic, Santa Rosa, Ca), Excluder (W.L.Gore & Associates, Flagstaff, Ar, USA), Zenith (William Cook, Europe, Bjaeverskov, Denmark) and, since 2002, the  EndoFit device (LeMaitre Vascular,  Burlington, Ma. USA). The aim of this study was to report on early and mid term results from our experience in 41 consecutive patients who received endovascular thoracic aortic repair with the EndoFit Thoracic stent graft system (LeMaitre Vascular).

Patients and Methods

A retrospective review of all patients who underwent endovascular grafting for thoracic aortic disease at our Institution from January 2002 to January 2007 using the EndoFit stent graft. Informed consent and institutional review board approval was obtained for all patients. Patient demographics, clinical and radiological data were obtained from hospital records.

The population consisted of 41 patients (33 men and 8 women) with an average age of 69.39.7 years (range 48 to 84 years ). Most patients had several risk factors and so were judged to have high open surgical risk (Table I). The indications for treatment were: progression of aneurysm size in 24 atherosclerotic aneurysms (58.54%) with a mean aneurysmal diameter of 7.191.48, acute contained aortic rupture in five patients (12.20%), (two of these presented aortoesophageal fistula as a complicating factor), recurrent pain and persistent hypertension or progression of aneurysm size despite maximal medical therapy in six patients with aortic dissection (14.63%), recurrent pain in four patients with penetrating atherosclerotic ulcers (9.76%), one post-traumatic pseudo-aneurysm ( 2.44%) and one patient with post-decoartation aneurysm (2.44%) (Table II).

All patients underwent spiral computed tomographic angiography (CTA) scanning of the chest, abdomen and pelvis with three-dimensional reconstruction before endovascular treatment to determine the location, length and diameter of the diseased aortic segment and the adequate size of distal vascular access. Stent graft dimensions were calculated from CTA images. Stent grafts were oversized by 10% to 20% of the landing zone diameter, but not in cases of aortic dissection were the oversizing did not exceed 10%. Proximal stent graft diameters ranged from 28mm to 42mm with only three of the grafts being implanted without free flow external fixation.

The anatomic suitability for endovascular grafting was determined by a proximal neck > 1.5 cm and a distal aortic neck > 2cm in length, so in three patients with a short proximal neck it was necessary to cover the origin of left subclavian artery. In these patients a careful examination of extra and intracranial circulation was performed.

General anaesthesia was used in 14 patients (34.15%) and spinal or epidural anaesthesia in 27 patients (65.85%). After surgical exposure, the common femoral artery was the most common site for insertion, but in three patients it was necessary to perform a retroperitoneal approach due to insufficient size of the femoral artery and in one patient a right subclavian approach was selected. A 6F pig tail was placed through the contralateral femoral artery to perform angiography for a correct graft deployment and to assess for endoleaks at the end of procedure. In the event of an endoleak, additional balloon dilatations with a special tri-lobe balloon (W.L. Gore, Associates Flagstaff, Az. USA) or an additional stent graft placement was performed. Intraoperative transesophageal echocardiography was routinely used in aortic dissection patients to position the stent graft into the true aortic lumen and to control the efficacy of the procedure. All patients received systemic heparin (70 IU/kg) and pharmacologic hypotension (systolic pick 80mmHg) was induced by anaesthesiologist at the time of graft deployment. More than one graft was used in 19 patients (46%), in these cases an overlap of at least 2-4 cm was performed.

EndoFit Thoracic Stent Graft System

The EndoFit Stent Graft (Fig.1A-B) is composed of self-expanding nitinol stents which are encapsulated in two layers of ePTFE (expanded polytetrafluoroethylene) fabric using a thermal process that avoids the need for fixation sutures. The Endofit sutureless design is an important benefit for optimal long term graft patency as the sutureless design eliminates the potential for suture breakage and graft material tearing. There is no interface of the metallic stents with blood or the aortic wall in the encapsulated section of the stent graft. There are two versions of the EndoFit stent graft, one with and one without an external fixation on the proximal end. The  external fixation is connected to the proximal end of the covered stent graft with two rods.

The introducer of the EndoFit stent graft is an hydrophilic 22F or 24F sheath with a tapered conical dilator over a 0.035” stiff guide wire. Stent graft loaded cartridges are also available which can be deployed through the hydrophilic introducer sheath (Fig.2). The EndoFit cartridge stent graft system is very useful as it is available in many sizes and takes up little space. The introducer system is also available with the loaded stent graft.

Platinum markers allow for precise placement and clear visualization of grafts. Deployment is effected using the push and pull technique. We recommend deploying the stent graft by using one smooth motion as spasmic movement could affect the accuracy outcome of the graft deployment.

The Endofit stent graft is currently available in the following sizes: proximal and distal graft diameter of 30–42 mm and a graft length of 12–24 cm. Proximal extenders and cuffs with and without external fixation are also available.

Also available is a true tapered stent graft (Fig. 1C) in sizes ranging from 42mm at proximal end down to 24 mm distally over a length of 24cm. We found this very usefull in treating chronic type B aortic dissection where a wide discrepancy of diameter from proximal to distal aorta is usually present with a very thick intimal flap. The ability to customize the EndoFit graft for tough torturous thoracic anatomy gives the surgeon or interventionalist a larger spectrum of patient population which can be treated endovascularly. We never used custom made stent-grafts although they are available on request. From October of 2006 the Flexi-Tip introducer tip was available with a shorter, more flexible tip designed to reduce trauma during the passage above the aortic arch. The FlexiTip design is very usefull in enabling exact and precise deployment  adjacent to the aortic valve (hybrid procedures).

Follow-up

In all patients a CT scan was routinely performed at one, six, twelve months and annually thereafter. A three month CT scan was done in patients with an endoleak. All significant events were recorded.

Statistical analysis

Each categorical factor is described as the number and percentage of patients. Continuous variables

are presented as mean ±standard deviation (SD) values in appropriation. The Kaplan Meier curves are used to plot overall and aneurysm related survivals over time. Statistical analysis was performed with SPSS 13.0 statistical software program.

Results

In all 41 patients the stent graft was deployed successfully (100% deployment success). There was no acute myocardial infarction, no need for dialysis, no respiratory failure, no graft migration, infection or fracture (Table III). We had only one spinal ischemia with consequent lower extremity weakness, no paraplegia was observed in our population. In three patients the proximal landing zone was judged too close to the left subclavian artery, so it was necessary to cover the left subclavian artery, but a left common carotid-left subclavian by-pass was considered unnecessary before stent graft deployment. No one of these patients developed  ischemic posterior circulation symptoms and left arm ischemia.

Patient Mortality and Morbidity

Aneurysm Related Mortality

Our definition for Aneurysm Related Mortality is a death which occurred before hospital discharge, or within 30 days of the procedure, or death due to aneurysm rupture. There were three in hospital deaths (7.32%). The first patient experienced an episode of hypotension and bradycardia after graft deployment, necessitating a temporary pacing and inotrope therapy. He died after two weeks for cardiac arrest in multisystem organ failure. The second and the third patients developed fatal cerebral vascular accident respectively one and two days after the procedure. There was only one real aneurysm related death 17 months after the procedure for aneurysm enlargement and rupture, the patient expired ten minutes after the arrival in emergency room.

Artery Access Complications

We had two artery access complications: one femoral artery dissection and one iliac artery rupture. Both were promptly surgically repaired without any consequence.

Endoleaks and Reinterventions

At discharge we recorded two proximal type I endoleaks, one spontaneously repaired thirty days post graft deployment and one required open surgical conversion after one year. One patient had a distal type Ib endoleak that self repaired after three months. During the follow up period we found two new type I endoleaks, one patient after twenty months and the other after 4 years. The endoleaks were treated with additional graft and both patients are in good health.

One other patient developed a retrograde type A dissection 15 months post procedure and underwent emergent replacement of the aortic valve and ascending aorta with an uneventful postoperative course.

Follow Up

Average follow up was 24.818 months. There were seven total deaths during the follow up period (17%), but only one was aneurysm related. The mean age was 76 years and all had several and severe co-morbidities including coronary artery disease (n=3), previous stroke (n=2), severe chronic obstructive disease (n=5), history of cancer (n=2), and concomitant abdominal aortic aneurysm repair (n=1).

The Kaplan –Meier survival curves (figure 3) shows patients freedom from all mortality and from aneurysm related mortality. At 2 years, the survival rates are 70% and 89% respectively.

In figure 4, Kaplan –Meier curve shows patients freedom from secondary endovascular or open surgery interventions.

Discussion

Endovascular stent graft treatment of patients with thoracic aortic disease was introduced  in 1992 at Stanford University and in 1994 Dake reported the results of their initial experience7. It was initially indicated for the treatment of descending thoracic aortic aneurysm with custom made devices for high surgical risk patients, subsequently also in patient with aortic dissection and penetrating atherosclerotic ulcers 8, 9.

Endovascular stent grafting in comparison to open surgery is a less invasive alternative in treatment of thoracic aortic disease, especially in elderly patients who very often present comorbidities including coronary and cerebrovascular disease, renal failure, and chronic obstructive pulmonary disease. In fact, in these patients, respiratory complications and severe renal failure represented the most frequent cause of morbidity and mortality, ranging from 5.3% to 33% and 2.5% to 13.3% respectively 10. Okita showed that the most severe complication after thoracic aortic surgery in septuagenarians and octogenarians was postoperative stroke which represented a strong incremental risk factor for not only the early death but also late death 11.

Another major complication in the open surgery of descending aorta is the spinal cord ischemia with consequent paraplegia. The main risk factors of spinal cord ischemia are perioperative hypotension, prior history of infrarenal abdominal aortic aneurysm repair, long segment coverage of the thoracic aorta compromising the intercostals arteries to the spinal cord. In fact, despite advances in surgical techniques, with reimplantation of intercostal arteries and shortened cross clamping times, this complication occurs in approximately 7-17% of cases 6 . In our series of 41 patients there was no need for dialysis, no respiratory failure, no acute myocardial infarction. No case of paraplegia was observed, we had only one spinal ischemia with consequent lower extremity weakness. This is an interesting result particularly if we compare this data to the spinal cord neurological events occurred in others endovascular studies 12, 13. Spinal cord protection maneuvers (spinal fluid drainage) were not used in these patients. According our experience, reducing the hypotension time during device deployment and then recovering a blood pressure ≥120/80 soon after it we believe that  it can reduce the spinal ischemic risk.

A careful study of an angiogram and CT scan is the first step to determine the anatomic suitability for endovascular grafting and to limit the vascular complications due to an insufficient size of distal vascular access. The EndoFit device has a hydrophilic introducer sheath of 22F and 24 F so, to limit vascular complications, we need a femoral artery with a diameter >7 or 8mm. In our study the common femoral artery was the most common site for access, but in three patients it was necessary to perform a retroperitoneal approach (12 mm Dacron conduit anastomosed end to site into the common iliac artery) due to the insufficient size of the femoral artery.  In one patient a right  subclavian approach was used due to the occlusion of the abdominal aorta. We had two vascular complications (4.8%), both were promptly surgically repaired. In other studies, 10-15% of vascular injuries are reported 14, 15. No in hospital conversion to open surgical repair was performed.

The anatomic suitability is defined by a proximal neck > 1.5cm and a distal aortic neck of > 2cm in length, so in three patients with a short proximal neck it was necessary to cover the origin of left subclavian artery. Many European centers have experienced few complications from covering the left subclavian artery16. It is our practice to carefully select  patients and reserve pre or perioperative carotid-subclavian by-pass for those in which it is necessary to preserve left internal mammary artery and/or left vertebral artery due to critical posterior cerebral circulation.

In hospital and 30 day mortality for emergent and non emergent cases treated by open surgical repair have been reported to range from 5.5% to 22.6%, up to 39% in patients with ruptured thoracic aortic aneurysm and acute aortic dissection17, 18, 19. The Eurostar and UK Thoracic Endograft registry reported a 30 day mortality of 9.3% 20. We recorded an in hospital and 30 day mortality of 7.32% (three patients); two of these were consequence of fatal strokes. They were old patients with several co-morbidities and severe atherosclerotic disease that could have been a predisposing condition to embolization days after deployment of the endoprosthesis.

We had a low incidence of endoleaks compared to other studies 21, 22, 23, 24. At discharge, only three patients (7.32%) showed endoleaks, one proximal type I endoleaks (Ia) and one distal (Ib), that were respectively self-repaired during one and three months of follow up. In fact some endoleaks have the tendency to thrombose spontaneously.  The third patient was discharged with a type Ia endoleak and required an open surgical conversion after the one year follow up period. . This patient had a type A aortic dissection and during the same admission to the hospital underwent to a combined open surgical repair of arch and ascending aorta with reimplantation of left carotid artery and endovascular treatment with EndoFit device of the descending aorta. Our strategy in patients with type B aortic dissection is to cover  the proximal entry point, and to stent the entire descending dissected aorta because aortic remodelling is induced by thrombosis of the false lumen 25,26.

During follow up, two patients (4.8%) developed type I endoleak, one after twenty months and the other after 4 years and both were treated with an additional endograft. No type II or III endoleaks were observed in our study population. It is our opinion that the advantages of the EndoFit  device as compared to the other stent grafts are the significant flexibility and  its tapered design. These features enable a customized and dedicated seating of the graft  in the thoracic aorta. It is also our experience that the flexibility of the introducer system minimizes vascular complications compared to the other devices (13) (14). Moreover the flexible construction of the introducer permits usage in sharply curved aortic anatomies without kinking and no stent fractures or collapse were recorded in our 41 patients (Fig.5). Our data are similar to those of Saratzis 27 and different from those reported by Melissano and Al. 28. We believe that the disappointing results of Melissano were the consequence of selecting patients with very critical anatomies in their initial learning curve using this device. Moreover, our patients were at high surgical risk; six were with aortic dissection, four with recurrent pain with penetrating atherosclerotic ulcers, and acute contained aortic rupture in five patients was present.  As in Saratzis study, we did not select patients, and we used the EndoFit device also to treat very difficult anatomies.

As with the Melissano and Saratzis studies, no spinal cord ischemia was observed in our study. One patient in our study was observed with lower extremity weakness after the procedure (paraparesis). There were no patients with paraplegia.. This is an interesting result particularly if we compare these data with others using different endovascular devices 12, 15, 20, 29.

The natural course of untreated thoracic aortic aneurysm is aneurysm rupture with a 2 years survival of 30% 30. A recent study reported a 3-years survival of 50% to 70% after surgery of thoracic aortic aneurysm and aortic dissection 31. All of our patients were not ideal for open surgery. In our patient population, the late mortality rate was 17% and depended largely on underlying pathologies.  In fact, only one was aneurysm related. Average follow up was 24.818 months. The Kaplan –Meier survival curves (figure 2) show patients freedom from all mortality and from aneurysm related mortality. At 2 years, the survival rates are 70% and 89% respectively.

Conclusion

Our results suggest that the EndoFit stent graft system gives to the endovascular specialist the possibility to treat even high surgical risk patient population with confidence and overall good results. This is largely due to the technical features of this device and availability of a wide range of  stent grafts for specific anatomies and also taking advantage from the flexibility of the EndoFit introducer sheath. We found few vascular access complications, no fractures or migration of the stent-graft in our follow up. With all variables considered, we can conclude that positive results have been obtained with the EndoFit stent graft. Longer follow up is needed to confirm these short and mid term results.

References

- Fabian TC, Richardson JD, Croce MA, Smith JS Jr, Rodman G Jr, Kearney PA, et al. Prospective study of blunt aortic injury: multicenter trial of the American Association for the Surgery of Trauma. J Trauma 1997; 42: 374-83.

- Clouse WD, Hallet JW Jr, Schaff HV, Gayari MM, Ilstrup DM, Melton LJ. Improved prognosis of thoracic aortic aneurysm: a population based study. JAMA 1998; 280: 1926-9.

- Svensson LG, Labib SB. Aortic dissection and aortic aneurysm surgery. Curr Opin Cardiol 1994; 191-9.

- Chiesa R., Melissano G., Civilini E., De Moura ML., Carozzo A., Zangrillo A. Ten years experience of thoracic and thoracoabdominal aortic aneurysm surgical repair: lessons learned. Ann. Vasc. Surg. 2004; 18: 514-20.

- Buffolo E., da Fonseca JHP, de Souza JAM, Alves CMR. Revolutionary treatment of aneurysm and dissections of descending aorta: the endovascular approach. Ann Thorac Surg 2002; 74: S1815-7.

- Attar S, Cardarelli MG, Downing SW et al. Traumatic aortic rupture: recent outcome with regard to neurologic deficit. Ann Thorac Surg 1999; 67: 959-65.

- Dake MD, Miller DC, Semba C, Mitchell RS, Walker PJ, Liddell RP. Transluminal placement of endovascular stent grafts for the treatment of descending thoracic aortic aneurysm. N Engl J Med 1994; 331: 1729-34.

- Dake MD, Miller DC, Mitchell RS, Semba CP, Moore KA, Sakai T. The first generation of endovascular stent graft for patients with aneurysm of the descending thoracic aorta. J Thorac Cardiovasc Surg 1998; 116: 689-704.

- Murgo S, Dussaussois L, Golzarian J, et al. Penetrating atherosclerotic ulcer of the descending thoracic aorta: treatment by endovascular graft. Cardiovasc Intervent Radiol 1998; 21: 454-8.

- Svensson LG, Crawford ES, Hess KR, Coselli JS, Safi HJ. Variables predictive of outcome in 832 patients undergoing repairs of the descending thoracic aorta. Chest 1993; 104: 1248-53.

- Okita Y., Takamoto S., Ando M., Morota T., Yamaki F., Kawashima Y., Nakajima N. Predictive factors for postoperative cerebral complications in patients with thoracic aortic aneurysm. Eur J Cardio-thorac Surg 1996; 10: 826-32.

- Grayson H.Wheatley III, Ahmet Tayfun Gurbuz, Julio A. Rodriguez-Lopez, Venkatesh G. Ramaiah, Dawn Olsen, James Williams, Edward B. Diethrich. Midterm outcome in 158 Consecutive Gore TAG Thoracic Endoprostheses: single center experience. Ann Thorac Surg 2006; 81: 1570-7.

- Cambria RP, Brewster DC, Lauterbach SR, et al. Evolving experience with thoracic aortic stent graft repair. J Vasc Surg 2002; 35: 1129-36.

- Ishida M, Kato N, Hirano T, Cheng SH, Shimono T, Takeda K. Endovascular stent graft treatment for thoracic aortic aneurysm: short to mid term results. J Vasc Interv Radiol 2004; 15: 361-7.

- Makaroun M.S., Dillavou E.D, Kee S.T, Sicard G., Chaikof E., Bavaria J, Wiliams D., Cambria R.P, Mitchell R.S. Endovascular treatment of thoracic aortic aneurysm: results of the phase II multicenter trial of the GORE TAG thoracic endoprosthesis. J Vasc Surg 2005; 41: 1-9.

- Gorich J, Asquan Y, Seifarth H, Kramer S, Kapfer X, Orend KH, et al. Initial experience with intentional stent graft coverage of the subclavian artery during endovascular thoracic aortic repairs. J Endovasc Ther 2002; 9: SII39-43.

- Galloway AC, Schwartz DS, Culliford AT, RibakoveGH, Grossi EA, Esposito RA, et al. Selective approach to descending thoracic aortic aneurysm repair: a ten years experience. Ann Thorac Surg 1996; 62: 1152-7.

- Glade GJ, Vahl AC, Wisselink W, Linsen MAM, Balm R. Mid Term survival and cost of treatment of patients with descending thoracic aortic aneurysm; endovascular vs open repair; a case control study. Eur J Vasc Endovascular Surg 2005; 29: 28-34.

- Fann JI, Smith JA, Miller DC, Mitchell RS, Moore KA, Grunkemeier G, et al. Surgical management of aortic dissection during a 30-year period. Circulation 1995; 92 (9 Suppl): II 113-21.

- Leurs LJ, Bell R, Degrieck Y, Thomas S, Hobo R, Lundbom J. Endovascular treatment of thoracic aortic diseases: combined experience from the EUROSTAR and United Kingdom Thoracic Endograft registries. J Vasc Surg 2004; 40: 670-80.

- Criado FJ, Abul Khoudoud OR, Domer GS, et al. Endovascular repair of the thoracic aorta: lesson learned. Ann Thorac Surg 2005; 80:857-63.

- Katzen BT, Dake MD, MacLean AA, et al. Endovascular repair of abdominal and thoracic aortic aneurysms. Circulation 2005; 112: 1663-75.

- Patel H.J, Williams DM, Upchurch Jr G.R., et al. Long Term results from a 12 year experience with endovascular therapy for thoracic aortic disease. Ann Thorac Surg 2006; 82: 2147-53.

- Marcheix B., Dambrin C., Bolduc J.P., et al. Midterm results of endovascular treatment of atherosclerotic aneurysm of the descending thoracic aorta. J Thorac Cardiovasc Surg 2006; 132: 1030-6.

- Nienaber CA, Fattori R, Lund G, et al. Non surgical reconstruction of thoracic aortic dissection by stent graft placement. N Engl J Med 1999; 340: 1539-1545.

- Dake MD, Kato N, Michell RS, et al. Endovascular stent graft placement for the treatment of acute aortic dissection. N Engl J Med. 1999; 340: 1546-1552.

- Melissano G., Tshomba Y., Civilini E., Chiesa R. Disappointing results with a new commercially available thoracic endograft. J Vasc Surg 2004; 39: 124-30.

- Saratzis N., Saratzis A., Melas N., Ginis G., Lioupis A., Lykopoulos D., Lazaridis J., Kiskinis Dimitrios. Endovascular treatment of descending thoracic aortic aneurysm with endofit stent graft. Cardiovasc Intervent Radiol 2007 30: 177-81.

- Najibi S, Terramani TT, Weiss VJ, et al. Endoluminal versus open treatment of descending thoracic aortic aneurysm. J Vasc Surg 2002; 36: 732-7.

- Crawford ES, Hess KR, Cohen ES et al. Ruptured aneurysm of the descending thoracic and thoracoabdominal aorta. Analysis according to size and treatment. Ann Surg 1991; 213: 417-25.

- Brandt M, Hussel K, Walluscheck KP, Boning A, Rahimi A, Cremer J. Early and long term results of replacement of the descending aorta. Eur J Vasc Endovasc Surg 2005; 30: 365-9.